## Slide 1: Smart noise judgment for quality test of X-Ray tube SXVT X-Ray Tube Project T CED INW-CN | April,2022

## Slide 2
Smart noise judgment for quality test of X-Ray tube
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 2

## Slide 3: Introduction
Background
- SXVT Tube is the important part of X-ray machine.
- Typical failure types are Noise, Bearing(Anode) stuck, Shaft-broken.
- Currently test in noise test room by sound-meter, record an observation with single number value.
- Need an objective standard to judge the quality of the X-Ray Tube.
XTA
Target
- 记录/存储 当前产线测试房的 旋转管的噪声测试数据.
- 探索并分析 XTA的噪声数据，确定信号特征.
- 构建实时分析模型，对测试的XTA 提供客观噪声合格与否的判断结果.
- 产品噪声质量追溯应用.
Note: Picture only for reference
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 3

## Slide 4: Feasibility Analysis
Data Features
Data Acquisition
Insight of noise data features (based on sound files)
Spectrogram analysis of sound signals (analogue data, incl phone’s recording, NTi and AiHua Sound Meter)
Difference of frequency spectrum energy
Difference stability in data
good
abnormal
high
Explore DA Solution for on-site sound data
Considering factors: ensure real-time, sampling frequency, effort and cost.
Nti XL2
Aihua AWA6228+
Digital (discrete) data is the basis for features extraction
Amplitude
Distribution
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 4

## Slide 5: Solution Overview
Field Level*
Services/Apps
Edge Level
噪声质量测试 客户端
* Take one test room as example in the architecture
Test Room
Test Bench
Mic
Sound Meter
AWA6228+
Server / PC
数据收集 & 边缘计算
时序数据存储
1
Product ID
2
串口控制及通信
基于Client交互控制，准确记录各转向的测试数据（消除脏数据干扰）
展示当前测试的数据
分析结果（Normal/Abnormal）提示
3
配置及数据获取
统计与分析
事件分析/判断
4
5
质量归因分析 应用
6
On-Site
Web
Service
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 5
支持用户基于产品编号和时间查询测试数据，和客户端分析结果
实现对时间周期内的产品质量的统计和分析展示

## Slide 6: Prototype 1.0 构建完成 现场Client端的 功能部分
Field Level*
Services/Apps
Edge Level
噪声质量测试 客户端
* Take one test room as example in the architecture
Test Room
Test Bench
Mic
Sound Meter
AWA6228+
Server / PC
数据收集 & 边缘计算
时序数据存储
1
Product ID
2
串口控制及通信
基于Client交互控制，准确记录各转向的测试数据（消除脏数据干扰）
展示当前测试的数据
分析结果（Normal/Abnormal）提示
3
配置及数据获取
4.1
On-Site
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 6
Target：
实现现场的有效测试数据的存储
Cost：
￥253,000 （先完成一个静音房），
后续可以+2w 完成另一个静音房的集成

## Slide 7: Prototype 2.0 现场Client端 集成数据分析
Edge Level
噪声质量测试 客户端
* Take one test room as example in the architecture
Server / PC
数据收集 & 边缘计算
时序数据存储
串口控制及通信
基于Client交互控制，准确记录各转向的测试数据（消除脏数据干扰）
展示当前测试的数据
分析结果（Normal/Abnormal）提示
配置及数据获取
4.2
On-Site
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 7
统计与分析*
事件分析/判断
Services/Apps
5
Target：
实现信号特征分析，提供客观的基于数据的分析判定结果
*：建议基于1/3 OCT 输出各频带值进行分析
Cost：
￥234,000 , Ray和SDR类型产品

## Slide 8: Prototype 3.0 产品噪声质量追溯 应用
Services/Apps
Edge Level
* Take one test room as example in the architecture
Server / PC
数据收集 & 边缘计算
时序数据存储
配置及数据获取
统计与分析
事件分析/判断
质量追溯 应用
支持用户基于产品编号和时间查询测试数据，和客户端分析结果
实现对时间周期内的产品质量的统计和分析展示
6
Web
Service
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 8
Target：
从数据的角度，实现产品数据追溯，统计噪声质量记录，直观展示
Cost：
预估￥409,500

## Slide 9
Support Needed and Risk
Restricted | © Siemens 2022 | YU HE | T CED INW-CN
Page 9
Support Needed:
    IT support: establish demo IT environment, including local network and internet.
    Hardware support: sound-meters and micro-phone ready, RS232 line, ToUSB connector.
    Software support: sound-meter’s functions package.
    X-Ray tube expert support: explain tube’s real scenario and the discuss failure mode, like good / bad noise of X-Ray products.
    Engineer support: support on-site test and help to provide test data and information, especially the failure data in multi-types.
    Requirement definition and discussion: from business view define requirement and confirm on solution workflow and UI design.
Risk:
    Sound-meter’s functions package data quality and impacted by environment.
    Less of different product types whole status mode and data.
    Requirement of service-design still need clarify.

## Slide 10
Questions?
Answers!

## Slide 11: Contact
Published by Siemens
T CED INW-CN
Expert
  YU He
200, Linghu Avenue 214135, Wuxi China
Mobile: +86 18921506861
E-mail: yu-he@siemens.com